From:	VASAGAM Jehan -PICKERING

Sent:	Tuesday, November 29, 2022 2:27 PM

To:	WATLING John -ENTENG

Cc:	ROQUE Clinton -PICKERING

Subject:	RE: Emonitor Load/Unload Screen Refresh

Categories:	Red Category

Hi John,

It seems like the operating system may be the reason.. the dataPAC 1500 was on unknown aswell and didn’t have the refresh problem.

Here is the details for A side

B side

Jehan Vasagam

Predictive Maintenance

Ontario Power Generation

(905)839-1151  ext 4401| Jehan.Vasagam@opg.com

See how we are powering the future of Ontario at opg.com

_____________________________________________
From: WATLING John -ENTENG <john.watling@opg.com> 
Sent: Tuesday, November 29, 2022 7:48 AM
To: VASAGAM Jehan -PICKERING <jehan.vasagam@opg.com>
Cc: ROQUE Clinton -PICKERING <clinton.roque@opg.com>
Subject: FW: Emonitor Load/Unload Screen Refresh

Hi Jehan,

To confirm – you are getting the Emonitor Load/Unload Screen Refresh occurring with the Enpac 2500 Card?

John Watling, P.Eng.

Senior Technical Engineer | Mechanical Thermal Components

<< OLE Object: Picture (Device Independent Bitmap) >>

(905) 839-1151 ext 5464 TUE/WED/THU | (647) 988-8778 MON/FRI | john.watling@opg.com

1675 Montgomery Park Road, Pickering, ON  L1V 2R5

See how we are powering the future of Ontario at opg.com

_____________________________________________
From: WATLING John -ENTENG 
Sent: Tuesday, May 31, 2022 3:40 PM
To: ROQUE Clinton -PICKERING <clinton.roque@opg.com>
Cc: ALI Andrew -ENTENG <andrew.ali@opg.com>; MCNEIL Cecil -C I O <cecil.mcneil@opg.com>
Subject: Emonitor Load/Unload Screen Refresh

Hi Clinton,

The constant Load/Unload screen refresh does not occur for other data collectors (i.e. for example the dataPAC 1500);

<< OLE Object: Picture (Device Independent Bitmap) >>

Nor does it happen for other data collector cards, like the dPAC 1500 card);

<< OLE Object: Picture (Device Independent Bitmap) >>

Nor does it happen for the Dynamix USB;

<< OLE Object: Picture (Device Independent Bitmap) >>

The screen refreshing only seems to occur with the Dynamix card arrangement, so is most likely a bug within Emonitor.

I will reach out to the Rockwell bulletin board for support.

John Watling P.Eng.

Senior Technical Engineer | Mechanical Thermal Components

<< OLE Object: Picture (Device Independent Bitmap) >>

(905-839-1151 ext 5464 TUE/WED/THU| 647-988-8778 MON/FRI | john.watling@opg.com

1675 Montgomery Park Road, Pickering, Ontario L1V 2R5

See how we are powering the future of Ontario at opg.com